## Slide 1
2023- Boston: Interferometric nanopositioner
He-Ne laser
Detector bench
Mobile corner cube
Fixed corner cube
Air-bearing stage
Voice coil actuator
Turning mirror
Capacitance probe

## Slide 2
2023 Sponsors
Physik Instrumente gmbh
VDL Enabling Technologies Group
RepairFAQ.org

## Slide 3
Equipment support
He-Ne laser
Mechblocks
Machining support
Capacitance probe
linescale
Voice coil
Air-bearing stage

## Slide 4
Interferometer layout

## Slide 5
Sam Goldwasser to share basics of Homodyne interferometry shortly
z
PBS
x
y
Laser is rotated in z such that it‘s linear polarization state is 45° (as shown in left) before entering the polarizing beam splitter (PBS)
Linearly polarized laser
(X and Y components have equal magnitude (intensity) and zero phase offset)
https://commons.wikimedia.org/wiki/File:Linearly_polarized_wave_-_45_degrees.gif
https://commons.wikimedia.org/wiki/File:Circular_LH_polarization.gif
P
50:50 beam splitter
P
S
S
Corner cube
QWP
S
P
CP2
QWP
Linear Polarizer
Photodiode 2
QWP
Linear Polarizer
Circular Polarizer, CP1
Photodiode 1
Circularized beams interfere
QWP
Corner cube
S
P
No interference occur here because the reflected waves oscillate in orthogonal planes
P
S
When the rotation angle of the Circular Polarizers differ by 45° and any of the corner cubes move,
Photodiode 1
Photodiode 2
Intensity
Time
FRINGE

## Slide 6
Assemble and align
the interferometer with the preassembled stage
He-Ne laser
Turning mirror
Detector bench
Fixed corner cube
Mobile corner cube

## Slide 7
Laser mounts (2x) with M4 setscrews
Laser
PH1.5 optical post holders (2x) attached to the optical breadboard using ¼-20 setscrews
TR50 optical posts (2x)
Base clamp (BA1SL )
¼-20 L0.75 in. bolt and washer
optical post (TR50)
optical post holder (PH50)
2x1 Mechblock
Laser mounting & steering
Copper coated turning mirror
M4 half round washers
M4 L12 screw
M4 L10 screw

## Slide 8
Moving corner cube
Use base clamp (BA1S ) for lateral adjustment
¼-20 L0.75 in. bolt and washer
Fixed corner cube
optical post (TR50) for vertical adjustment
optical post holder (PH50)
optical post holder (PH30)
optical post (TR20) inside
M6 L16 bolt and washer
Corner cubes
¼ wave plates will need be attached on the face of the corner cubes

## Slide 9
Polarized beam splitter on a tip tilt mount

## Slide 10
Fixed corner cube
Moving corner cube
Laser source
Detector
Moving carriage
Fixed guideway
beamsplitter
Fixed corner cube
Moving corner cube
Laser source
Detector
If stage motion axis is not inline with the laser beam axis
This will cause the entry and exit beams on the moving corner cube to translate (‘shear’)
Therefore, the exit beam will not interfere with the beam from the fixed corner cube
beamsplitter

## Slide 11
Fixed corner cube
Moving corner cube
Laser source
Detector
One can rotate the stage, or the whole intererometer asssembly
beamsplitter
Or simply rotate the beamsplitter
Fixed corner cube
Moving corner cube
Laser source
Detector
beamsplitter
Therefore, the tip tilt adjustment

## Slide 12
Polarized beam splitter
50:50 beam splitter
Photodetector 1
Photodetector 2
From mobile corner cube
From fixed corner cube
Circular Polarizer 1
Circular Polarizer 2
Detector bench
Optical post (TR50) and holder (PH30)
Tip-tilt kinematic adjuster (KM05/M)
¼-20 screws to mount to the optical breadboard
Note that photodetector PCBs are mounted so that the cables come from the same side

## Slide 13
Identifying magnetically held polarization optics
¼ wave plates – yellow chrome plated steel washer, 6 mm aperture
Circular polarizers – Zinc plated steel washer, 6 mm aperture
Linear polarizers – Zinc plated steel washer, 3 mm aperture
https://www.youtube.com/watch?v=qwOXKlQ_gFo&t=296s

## Slide 14
M4 threaded inserts
Loosen the M4 L20 screws to adjust the beam splitter base clamps laterally
Base clamp
Base clamp
M4 threaded inserts
M4 half round washer
M4 threaded inserts
M4 half round washer
M4 threaded inserts
M4 half round washer
¼-20 L0.75 in.
5x5 Mechblock
5x1 Mechblock
Two 5x0.25 Mechblocks
Optical post holder (PH30)

## Slide 15
M4 L16 screw
M4 L12 screw
M4 half round washer
Loosen the M4 L16 screws to adjust the photodetectors laterally
50:50 beam splitter
M4 L10 screw
M4 half round washer

## Slide 16
Capacitance probe holder
M4 L12 screw
M4 half round washer
Two ¼-20 screws to mount to the optical breadboard
M4 threaded insert
M4 L12 screw and washer
probe
Clamp
C11-1.5-9.5-N

## Slide 17
Power supply for the 16-bit DAQ -voice coil amplifier box
-15 V
+15 V
GND

## Slide 18
Stage voice coil voice coil Jack plugs connects to AMP 1
PCB connector block for encoders, limit switches, and DAQ communication. Goes to connector A of myRIO
Lion precision capacitance gage
BNC cable connects to ADC 1 of the DAQ unit
Optical limit switches from the stage connects to 6-pin Molex connector
PCB for the photodetectors
Shielded ethernet cable
DB15 connector to read the air bearing stage linescale
Hardware communication
DAQ/AMP unit
15 V power supply
A
B
NI-myRIO microcontroller
A
B

## Slide 19
Three sensors
Interferometer
Linescale
Capacitance probe

## Slide 20
Sensor advantages
Range
Resolution
Speed
Interferometer
Linescale
Capacitance probe
For stabilized He-Ne laser,
Inherent calibration:
Each fringe corresponds to a displacement of
Range in meters
3 nm resolution at 0.25 mm/s
50 µm Range
1.8 nm resolution at
15 kHz bandwidth
10 mm Range
70 nm resolution
10’s of mm/s velocity

## Slide 21
Sensing Hardware Specifications
Interferometer resolution limitation due to myRIO’s B side ADC: 5V range, 12-bit resolution, 250 ksps sample rate when reading two channels
If the fringe intensity is 0.5 VPV, interferometer resolution is
RT LabVIEW program provided to calculate displacement can run at 120 us (8.33 kHz).
  https://www.youtube.com/channel/UCEntCZGnnCwgXbp8-1AV49w
  https://github.com/ASPE-student-challenge/2023
Say the air-bearing stage travels at 0.25 mm/s, 790 fringes would be produced in a second (0.25 mm/316.4 nm).
To sample 100 datapoints per fringe (3.16 nm resolution), one needs to sample at 7.9 kHz
For He-Ne laser, ideally
Each fringe (one complete rotation of the circle) corresponds to a displacement of
Photodiode 1
Photodiode 2
Intensity
Time
VPV

## Slide 22
Linescale: 70 nm resolution
LabVIEW FPGA program reads it at 8 ticks (40 MHz/8 = 5 MHz)
For example, if the air-bearing stage travels at 10 mm/s, would produce only 143k cycles of line scale output per second (10 mm/70 nm), and can be comfortably read with the 5 MHz sample rate.
Sensing Hardware Specifications
Capacitance probe
Range: 50 µm (+/- 10 V), 2.5 µm/V sensitivity, read the datasheet for non-linearity error
Resolution: 0.4 nm at 100 Hz bandwidth, 1.8 nm at 15 kHz bandwidth
https://www.lionprecision.com/wp-content/uploads/2022/07/Brochure-CPL230_350.pdf
resolution limit of the 16-bit ADC to read the probe is

## Slide 23
Before the preliminary presentation on Oct 20th
Watch the YouTube videos and use the solid model assembly (Github) to assemble the interferometer.
  https://www.youtube.com/playlist?list=PLy1-ZnpGPTPVfVI97nkTYyfq9PLcodRp7
  https://github.com/ASPE-student-challenge/2023
Install LabVIEW (instructions at the end of the slides), use the sample project in Github. A sample program controls the stage using linescale feedback, uses interferometer to monitor the motion
Modify the LabVIEW program to feedback control the stage using the interferometer
Think about using all three sensors (interferometer, line scale, and capacitance probe) in combination for feedback, and as diagnostics
Tune the PID controller to improve the positioning performance. Lei Zhou, University of Wisconsin, Madison is expected to provide a tutorial in one of the upcoming weeks
Prepare the error budget. Stuart Smith, University of North Carolina at Charlotte is expected to provide a tutorial in one of the upcoming weeks
Post your questions on Reddit https://www.reddit.com/r/aspestudentchallenge/

## Slide 24
ASK FOR HELP
aspe.challengecommittee@gmail.com

## Slide 25
Deliverables for the preliminary presentation (200 points)
The meeting will be held via Zoom on Oct 20, 2023, and the schedule for each team will be emailed for their 4-minute presentation. Committee will provide live feedback and address questions of individual teams.
Teams are required to present their progress in getting the stage and interferometer operational (metrology and control efforts ) with well labelled images
10 second interferometer stability sampled at 100 Hz without air supply to stage
Report 50 nm, 5 um, 500 um, 5 mm step responses with 5 s dwell time sampled at 100 Hz using interferometer or encoder feedback. Report the reading of the sensor not used for feedback (line-scale or interferometer).
Preliminary error budget
50 nm
5 s
5 µm
5 mm
5 s
500 µm

## Slide 26
Conference deliverables
Position control the stage for a travel profile provided by the committee. Details will be revealed soon.
By Nov 5th midnight, submit a strictly two-page report (maximum of up to 4 figures) including the error budget, challenges overcame, and novel implementations.
3. You can update the error budget by Nov 13th midnight after assessing the challenge environment at the conference.
4. Presentation of metrology and control efforts, novel implementation, and Q&A.

## Slide 27
Key dates
Oct 20th

## Slide 28
https://github.com/ASPE-student-challenge/2023
https://www.youtube.com/channel/UCEntCZGnnCwgXbp8-1AV49w
https://www.reddit.com/r/aspestudentchallenge/
Resources
aspe.challengecommittee@gmail.com

## Slide 29
LabVIEW installation
A list of required software for myRIO can be found here, and LabVIEW myRIO software bundle contains all the required software.
To check the LabVIEW software and version installed on your PC and myRIO, open the NI MAX program on your PC. For software installed on the PC, navigate to My System > Software.
For software installed on the myRIO, ensure that the myRIO is connected to the PC, and navigate to Remote Systems > name of your myRIO > Software.
  On your PC, in general
LabVIEW 2019 with FPGA and Real Time Modules
NI-RIO 19
Example list of NI software downloaded on the PC as shown in MAX. Indicated are the necessary modules.
  On myRIO, in general
LabVIEW Real-Time 19.0.09
NI myRIO 19.0
List of NI software installed on the myRIO used during testing as shown in MAX. Some but possibly not all of these modules are necessary.

## Slide 30
  Installing software on the myRIO
Connect the myRIO to the PC. Open the NI MAX program. Navigate to Remote Systems > name of your myRIO > Software. Right click on Software and click Add/Remove Software.
Next, enter the device password for the admin user (likely blank) and select “NI myRIO 19.0” under the “LabVIEW Real-Time 19.0.0” tab. Click Next.
Select all the software set add-ons, click Next.
Click Next on the summary page and the software will be installed on the myRIO. The device may restart several times; expect the process to take several minutes.

## Slide 31
Organizing committee
Luis A. Aguirre – University of Texas, Austin
Kumar Arumugam – NIST
Raymond C. Cady – Corning Inc.
Jacob Cole – IST precision
Sam Goldwasser – RepairFAQ.org
Liam G Connolly – NIST
Tim Dalrymple
Chunjie Fan – Corning Inc.
Ping Guo – Northwestern university
Mark T. Kosmowski – Electro-Scientific Industries, Inc.
Todd Noste – Lawrence Livermore national laboratory
Charles Perera – NASA, Goddard space flight center
Jesse Redford - University of North Carolina at Charlotte
Nilabh Roy – Canon Nanotechnologies
Stuart Smith – University of North Carolina at Charlotte
Lei Zhou - University of Texas, Austin
With contributions and consultant input from
Dhanooj Bobba - University of North Carolina at Charlotte
Tyler Hamer – MIT Lincoln laboratory
Hilary Johnson – Lawrence Livermore national laboratory
Sam Ludwig– Lawrence Livermore national laboratory
Thapelo Mametja - National Metrology Institute of South Africa
Parth Pandya - University of Texas, Austin
Kent Queensland – Lion precision
Senajith Bandara Rekawa – Lawrence Berkeley Laboratories